## Slide 1: Qualifying Surface and Device Performance Through Roughness Quantification Samuel.Lesko@bruker.com

## Slide 2: Welcome to the webinar
For best audio experience, everyone is on mute
For question, use the Chat
Questions will be answered at end of webinar
Webinar is recorded and will be available offline
  Feel free to share & forward to colleagues

## Slide 3: What you will learn today
How to select best means to measure roughness for your applications
When to decide between profile & areal roughness measurement
Which benefits areal S parameters from ISO 25178 bring
Which filter cut-off to chose for meaningful roughness
What the most relevant parameter(s) is (are) versus specific applications

## Slide 4
Forewords

## Slide 5: Background Part from Bruker – Nano Surfaces division
Atomic Force Microscopy
3D Optical Microscopy
WLI & Focus Variation from Bruker Alicona
Stylus Profilometry
Universal Tribology
& Mechanical Testers
MultiPhoton, Super Resolution
Fluorescence Microscopy
Hysitron Nanomechanical Test Instruments
Bruker Nano Surfaces Division Pioneers in Microscopy, Metrology & Mechanical Testing

## Slide 6: Roughness measurement Needs in Industry
Quantify Wear…
  Failure investigation
  Design for longer life
  Life time warranty
Quantify Texture…
  Converge to efficiency
  Environmental: low noise, leak, emission
  Aesthetics: vibration, appearance
Quantify Manufacture Process…
  Monitor production
  Production cost
  Choose supplier

## Slide 7: Roughness measurement Range

## Slide 8: Roughness measurement Profile vs. Area
Roughness has been measured for decades via profile cross section
What industrial surfaces lend themselves to area inspection?
CNC surfaces (turning, grinding, milling…)
Single Profile
Area measurement
Textured Surfaces (honing, laser,…)
Multi - Single Profiles
Area measurement
Random Surfaces (defect, additive manufacturing, abrasive…)
Single Profile
Area measurement

## Slide 9: Roughness measurement Norms
Raw surface
Primitive Surface
Low pass / noise & spike removal
S-F Surface
Shape/form removal
S-L Surface
High Pass for roughness
Definition of area
Extraction S parameters
ISO 4287 & ISO 3274: cut-off defined by norm
ISO 25178: cut-off is defined by user

## Slide 10: Roughness measurement Potential pitfalls
1. Does instrument have proper lateral resolution?
2. Does instrument have adequate vertical resolution?
3. Does instrument have proper transfer function ?

## Slide 11
Technologies to measure roughness

## Slide 12: Roughness measurement Which system to select?
White Light Interferometry
Area
Sylus
Profile
SEM
NO TOPO
Height
Lateral resolution
Focus Variation
Area
Profile
Lateral details < 0.5µm
Thin film < 0.1µm
Textured / Slope
Vertical Precision < 0.1µm

## Slide 13: White Light Interferometry Principle
Pixel Intensity
Vertical scan
Best contrast
Camera
Moiré présence
Notes: FV is not new.

## Slide 14: White Light Interferometry Vertical resolution in all situations
Fold mirror for vertical wall charterization
Sa = 0.1µm!!!
No other optical profiling technique combines vertical resolution over large areas
Sa = 0.03 µm !!!

## Slide 15: White Light Interferometry Large dynamic & transfer function
Sa= 2.85µm – c > 0.8mm
Sa= 1.06µm – 0.08mm< <0.8mm
Sa=0.25µm – 0.008mm< <0.08mm
Surface component
Full vertical range: 10µm
True vertical resolution: 1nm
Unique vertical sensitivity versus all other non-contact profiling techniques
2 images
4.02 x 1.74 mm²
3.6µm lateral resolution

## Slide 16
Beyond mean roughness

## Slide 17: Roughness measurement Why Ra or Sa are not enough?
Same Ra
Same Sa

## Slide 18
Case study

## Slide 19: GAR Strip Corrosion Measurements How top choose cut-off?
Examined various surface finishes for rate of corrosion over longer time period
Similar Trends on increasing roughness and randomness of the surface
Rougher surfaces corroded to greater degree
Same
Different
Power Spectrum Density gives indication on which cut-off to select
High
Low
Peaks
Valleys
Notes: In a similar corrosion study as the last example, a steel surface is studied.
Surface finish is quantified.
GAR is the commercial name for a test strip of steel with difference textured surfaces and nominal roughnesses.

## Slide 20: Reflectivity efficiency Al coated mirror
Power Spectrum density emphasizes which spatial frequencies are the most important
In this case, intermediate spatial frequencies are key for appearance

## Slide 21: Reflectivity efficiency Al coated mirror
Proper spatial filtering reveals through mean roughness Sa differences between mirrors
Increase of roughness amplitude also increases diffusion of light
Diffusion directly scopes with reflectivity efficiency
Sa = 13.7nm
Sa = 6.5 nm
Mate appearance
Shiny appearance

## Slide 22: Quantification of opacity Glass manufacturing
Roughness filtered | Sku
Rough Side | 40
Soft Side | 10 000
Mat Side | 1 000
Prism Side | 200
Kurtosis
Notes: Please use:
Boxes with GREY outlines for pictures and text
Boxes with RED outlines for text only.
FULL GREY boxes for conclusions only.

## Slide 23: Quantification of efficiency Solar Cell
23
Notes: 23

## Slide 24: Entry qualification Cap for ultra-sound sensor
Good performances
Issues with coupling
X4 difference
X6 difference
Notes: Please use:
Boxes with GREY outlines for pictures and text
Boxes with RED outlines for text only.
FULL GREY boxes for conclusions only.

## Slide 25: Wear assessment Cylinder – Functional parameters
Vmp | 25 x 10-6
Vvc | 8.0 x 10-4
Vvv | 11.0 x 10-5
New
Used
Vmp | 8.3 x 10-6
Vvc | 1.5 x 10-4
Vvv | 1.8 x 10-5
1 mm
Abott / Firestone / Bearing curve
Vmp: peak material volume
Vvc: core void volume
Vvv: lubricant reservoir volume
Unit: (µm3/µm²)
Risk of adhesive behavior with piston

## Slide 26: Quantification of gloss Metal Belt ring
BAD
GOOD
Glossy
Shiny

## Slide 27: Quantification of gloss Metal Belt ring
There is clear differentiation of good/bad glossy surfaces for belt. Average slope variation indicates how shinny surface is with clear quantification.
GOOD: shiny
BAD: too mate

## Slide 28: Finding root cause of issue Brake vibration
Sharp well-defined triangular shape allows proper grid on disk
Spatial frequency plays a role here
Sdq combines
  Amplitude
  Lateral spacing
Sa: In spec
Sdq: High
Sa: In spec
Sdq: Low
Quiet
Noisy

## Slide 29: Predictive maintenance Sealing on rotating shaft
New shaft
Sa | 368nm
Sds | 1130 /mm2
Ssc | 49 mm-1 (radius 20mm)
Sa | 769 nm
Sds | 247 /mm2
Ssc | 6 mm-1 (radius 166 mm)
Used shaft
600 µm

## Slide 30: Optimization of process 3D printing of PEEK material
Sal = 77µm
Sal = 88µm
Notes: Measure high lateral resolution along line

## Slide 31: S areal roughness parameters Link with functionality
Strong Correlation
Weak Correlation
No Correlation

## Slide 32: Conclusion
3D optical profilometry allows proper description of complex texture of surface by areal measurement
Through large field of field and use of new ISO roughness parameters, 3D optical profiler techniques are able to rank surface effiency as well as to explain the reasons for failure
Selection of optical & stylus technologies is application dependant
  Stylus:
    quick long single trace
    multi-thin transparent layers < 100nm
  Focus Variation: large field with roughnesss
  White Light Interferometry:
    large field of view with sub-µm roughness
    high lateral resolution

## Slide 34: Sensor sensitivity When roughness is not enough
Chart: Number Pits > 1µm depth
| Good1 | Good2 | Medium1 | Medium2 | Bad1 | Bad2 | Bad3 | Bad4
Number Valleys > 1µm | 105 | 119 | 151 | 128 | 212 | 211 | 231 | 224
+20%
+50%
Good
Bad

## Slide 35: Best Practices Roughness not always enough

## Slide 36: Std – Spatial S-Parameter cont.
The figure below shows an image that has surface lay in it. It lays on the left side of the Y axis so Std = -17 deg
When the surface is isotropic then there is no lay present, so Std will read as N/A.

## Slide 37: Machining Influence on Texture Example 3D Areal Images – NB Directionality
Flat lapping/Reaming
Grinding
Horizontal Milling
Vertical Milling
Turning

## Slide 38: S Parameters – Hybrid Parameters (Sdq, Ssc, Sdr)
Sa=994nm, Sdr- 1.6%
Sa=1173nm, Sdr- 0.5%
Sdr may further differentiate surfaces of similar amplitudes and average roughness. As indicated above, two surfaces with only a difference of 10% in Sa have over a 200% difference in Sdr. Typically Sdr will increase with the spatial intricacy of the texture whether or not Sa changes.
Sdr is useful in applications involving surface coatings and adhesion. Sdr and may also find relevance when considering surfaces used with lubricants and other fluids. Sdr may be related to the surface slopes and thus may also find application related to the manner in which light is scattered from a surface.

## Slide 39: 1D Roughness measurement On textured cylinder
Native
Corroded
Roughness filtered profiles – Gaussian filter – cut-off 0.8mm

## Slide 40: Reason for areal roughness
Gear edge has multiple high angles (90 degrees) and must be characterized in 3D
2D calculations cannot reliably determine the geometry or roughness
WARNING: multiple 2D lines to build « 3D » require extra attention on how lines are paste together

## Slide 41: Orthopedic – Mean Roughness Qualification of polishing process on knee prostheses
Raw Casted Finish step 1 Finish step 2 Final hand buff

## Slide 42: Clutch & coupler Bearing metrics Sbi, Sci, Svi
Sa = 1.2 um Sbi=0.70 Vvc/Sz=1.31 Vvv/Sz =0.12
Sa = 4.2um Sbi=1.43 Vvc/Sz=0.46 Vvv/Sz =0.23
Clutch = grip surface
Coupler = transferring motion

## Slide 43: Car boards Spotting the difference on large areas
Ghosted
Non-Ghosted
More micro-structures seen in valleys

## Slide 44: S parameters show how they are different